Pycnometer Standard Operating Procedure

- Record your name, date, and other relevant information in the Pycnometery Log Book.

- Plug in the pycnometer to a power outlet and attach the gas hose to a gas cylinder, either helium or nitrogen.

- Power “ON” and allow 10-15 minutes for the pressure transducer to warm up and stabilize.

- Adjust the tank pressure regulator to slightly above 20 PSIG. Pressures above 25 PSIG can damage the pressure transducer.

- Select the correct REFERENCE VOLUME for the sample cell to be used. See the front panel of the pycnometer for reference volume selection.

Note: Select the reference volume where the sample cell can be filled ¾ full of sample material. The three sizes from smallest to largest are micro, small, and large.

Note: The toggle valves are CLOSED when their handles are parallel to the cabinet face and OPEN when perpendicular to the cabinet.

- Fill the sample cup, insert it into the cell holder, and replace the cover.

- Turn the selector valve to “CELL” if it is not already.

- Open the “GAS OUT” toggle valve and turn the “GAS OUT RATE” control completely counter-clockwise. Wait for a stable reading.

- Close the “GAS OUT” toggle valve, turn the “GAS OUT RATE” needle valve clockwise until resistance is felt (not completely closed), and set the meter to zero using the knob next to the digital readout.

- Turn the selector valve to “REF”.

- Rotate the “GAS IN RATE“ needle valve clockwise until it is just slightly open.

- Open the “GAS IN” toggle valve, and pressurize as close to 17 PSIG as possible using the “GAS IN RATE” needle valve to control the rate of pressurization. Stop the flow by closing the “GAS IN” toggle valve.

- Record the display reading after it has stabilized. This value is P1.

- Slowly turn the selector valve to “CELL”.

- Record the display reading after it has stabilized. This value is P2.

- Divide P1 by P2 and record this number.

- Vent the pressure slowly to prevent blowing powder out of the cell by opening the “GAS OUT” toggle valve with the “GAS OUT RATE” control slightly open.

- Repeat steps 9-16 until the ratio in step 15 is within ±0.005 for three consecutive measurements.

- Use the following equation to calculate the true powder volume.

Where VP is the volume of the powder, VC is the cell volume, and VR is the reference volume. VC and VP depend on the reference volume used are found under ”Pycnometer Calibration Values”.

- Remove the sample cup from the pycnometer and use an analytic balance to measure the mass of the sample cup with powder.

- Remove powder from the sample cup and measure the mass of the sample cup. The difference in masses is the mass of the powder

- Divide the mass by the volumes found in step 18 to find the density of the powder. The average of these measurements should be recorded as the density of the powder.

Pycnometer Calibration Values

LARGE CALIBRATION SPHERE	VCAL large= 56.559 cm3

SMALL CALIBRATION SPHERES (2)	VCAL small=2.145 cm3

LARGE SAMPLE CELL VOLUME	Vc large=151.7178 cm3

LARGE REFERENCE VOLUME	Vref large=90.2884 cm3

SMALL SAMPLE CELL VOLUME	Vc small=29.8415 cm3

SMALL REFERENCE VOLUME	Vref small=13.0142 cm3

MICRO CELL VOLUME	Vc micro=12.4670 cm3

MICRO REFERENCE VOLUME	Vref micro=7.0537 cm3

Calibration values measured on 07/12/2007 by Quantachrome Corporation.

Contact

Edison Buff 
Domestic Service Manger 
Quantachrome Instruments 
(561)  731-4999 
(561)  732-9888 fax 
ed.buff@quantachrome.com